RPW Operation Weekly Report – NECP
Outline
•	Performed Activities
•	Highlights – Anomalies
•	Upcoming Activities – Preparation
•	Science Performance
Performed Activities
•	RPW Short Functional Test (SFT) – 24-25th of February
•	HK TM data have been retrieved and processed. Data processing of other TM is on-going.
•	Preliminary analysis shows normal behavior of the instrument. RPW seems to work properly from a Cmd/Ctrl perspective.
•	Draft report has been distributed for internal review
•	Some low and medium severity anomalies have been observed
•	Calibration Rolls – 8h each day on 25-28th of February
•	Only data for the 25-26th have been retrieved and processed
•	Some low and medium anomalies have been observed
Highlights – Anomalies
•	Some anomalies have been observed during SFT :
•	Correctable EDAC errors (LE_DPU_AHB) - This was not expected but those correctable EDAC errors were already seen when the instrument is on and do not impact RPW behavior
•	Many errors while the compression is enabled (LE_DPU_COMP - CPKT_TOO_BIG)  Compressed packets are sometimes bigger than normal ones (in this case, these ones are transmitted). This has to be investigated.
•	One SW_PROCANOMALY error from TDS while in SBM2_ACQ  This has to be investigated
•	One TIMECODE_CTR error from THR at the end of each of its calibration  This has to be investigated
•	an error on W_BIA_M5V current @00:19:25 - The raw value of the TM is 1 LSB above the warning threshold while the bias current are set to the maximum allowed value on the redundant side. This is only a warning and do not lead to any reconfiguration of the instrument nonetheless the thresholds might have to be re-tuned
•	SCM temperature acquisition: On the LFR temperature curve the value of SCM temperature (blue curve) never goes under -53.8 °C which doesn’t look physical. The temperature curve is in accordance with raw telemetry values  This has to be investigated
•	An additional anomaly has been observed during first roll (not during SFT):
•	DAS buffer overflow – TM FIFO (?)  This has to be investigated
Upcoming Activities – Preparation (To be completed)
•	PDORs of the next week have been provided for IW-8 continuous activity.
Science Performance
•	Science telemetry analysis of deployment activities are still on-going, but strong interferences (80kHz and 120kHz) have been observed and seem to come from the platform (RW DC-DC converters and APR). These spurious are still present during the first hour of the calibration roll. If the EM noise from the S/C remain as it is now there will be a clear impact on the science associated the TNR-HFR unit. This impact cannot be mitigated by RPW. As it is the 120 kHz has also a strong impact on the TDS science. A patch of the TDS flight software will be needed to recover the science. 
AOB
•	EMC campaign configurations preparation is on-going
•	Inter-Instrument Communication campaign preparation is on-going